## Slide 1: ChemCam PLS1 Sub-Model Calibration

## Slide 2: Rationale for Sub-Models
We have had trouble coming up with single PLS1 models that can handle the full range of compositions
  This should not be surprising: samples with high concentration of an element may have different matrix effects than samples with low concentration
Expanding the database helped with this somewhat, but a single model still wasn’t enough
Instead, break the full composition range down into several sub-ranges, and develop a separate PLS model for each
  Optimize # of components and normalization separately

## Slide 3: Implementation
Use new Python code
  Can restrict database by composition on the fly
  Accepts input files to specify spectra to remove or keep, allowing outlier removal
  Automatically generates relevant plots
  Use the mlpy PLS algorithm, which gives identical results to our previous IDL PLS algorithm
Use database with Sam’s ratio-based calibration transfer

## Slide 4: Optimization
Use plots of:
  RMSECV (cross validation RMSE)
  RMSEC (RMSE of full model predicting the training set)
  RMSEP (RMSE of full model predicting the test set)
  RMSEP_cal (RMSE of full model predicting the 95A Mars cal targets)
Compare norm1 and norm3
Try to choose # of components corresponding to “elbow” in curves where possible
  Adjust for minima in cal target RMSEP, where relevant

## Slide 5: Predictions
Calculate a “full” prediction using a composition range of 0 to 100 wt.%
Calculate sub-model predictions (usually two or three sub-models)
Combine the predictions using simple logic:
  For example: If full model predicts <5 wt.%, use “low” submodel. If Full model predicts >15 wt.% use “high” submodel. Else use “mid” sub-model.